## Slide 1: Please delete this slide after reading through and before working on your slides
TIPS
Content provided in this lesson does not replace the i-STAT User Guides, System Manual or Cartridge Test and Information (CTI) sheets or Instructions for Use (IFU).
This power point may be modified to represent products within the i-STAT System currently in use within your facility.
Once the resource has been modified, customers will then assume responsibility for the content.*
Abbott will continue to provide and support the original resources made available on the www.globalpointofcare.abbott website.*
*Check the website regularly to ensure that you have the most up-to-date content.

## Slide 2: i-STAT 1 Orientation
i-STAT LEARNING SYSTEM | i-STAT 1 SYSTEM

## Slide 3: Learning Objectives
i-STAT LEARNING SYSTEM / system orientation
  i-STAT 1 System Components
  Anatomy of the i-STAT 1 Analyzer
  Basic Operation Process
  Cleaning & Disinfecting
Understand the following:
components
ANATOMY
OPERATION
CLEANING

## Slide 4: i-STAT 1 System Components
i-stat 1 learning system | i-stat 1 orientation

## Slide 5: System Components
i-STAT 1 ANALYZER
  Used to perform blood analysis testing with i-STAT cartridges, review test results, and conduct quality control (QC) testing
  Optional ability to transmit results via the i-STAT 1 Downloader/Recharger
COMPONENTS
ANATOMY
OPERATION
CLEANING
DOWNLOADER
Offers option to transmit results
i-STAT 1 ANALYZER

## Slide 6: System Components – cont’d
i-STAT 1 WIRELESS ANALYZER
  Used to perform blood analysis testing with i-STAT cartridges, review test results, and conduct quality control (QC) testing
  Wireless ability to transmit results via the built-in wi-fi radio
  Optional ability to transmit results via the i-STAT 1 Downloader/Recharger
DOWNLOADER
offers extra option for transmitting results
i-STAT 1 WIRELESS ANALYZER
COMPONENTS
ANATOMY
OPERATION
CLEANING
WIRELESS TRANSMISSION

## Slide 7: System Components – cont’d
i-STAT TEST CARTRIDGES
  Unit-use design
  Contains sensors and reagents for testing patient samples and quality control fluids.
COMPONENTS
ANATOMY
OPERATION
CLEANING
i-STAT 1 DOWNLOADER/RECHARGER
  Used to perform analyzer software updates.
  Recharges the i-STAT 1 9-Volt NiMH Rechargeable Battery
  Transmits results via wired network connection (optional)

## Slide 8: System Components – cont’d
DISPOSABLE BATTERIES & THE i-STAT 1 BATTERY CARRIER
  The analyzer requires two Ultralife 9-Volt lithium batteries for use with the i-STAT 1 Battery Carrier
CAUTIONS
Use only an Ultralife 9-volt lithium battery purchased from Abbott (List Number: 06F21-26).
The Ultralife 9-volt lithium battery has a safety feature that provides protection preventing the i-STAT 1 Analyzer from overheating due to component failure within the analyzer circuitry.
BATTERY CARRIER
9-VOLT LITHIUM
COMPONENTS
ANATOMY
OPERATION
CLEANING

## Slide 9: System Components – cont’d
i-STAT 1 9-VOLT NiMH RECHARGEABLE BATTERY (OPTIONAL)
  Alternate power source that can be charged using the i-STAT 1 Downloader/Recharger
CAUTIONS
Use only a rechargeable battery pack purchased from Abbott
Rechargeable battery packs not recommended by Abbott nor purchased from Abbott may be susceptible to overheating and could lead to a potential fire or burn hazard
COMPONENTS
ANATOMY
OPERATION
CLEANING
9-VOLT NiMH

## Slide 10: System Components
i-STAT 1 ELECTRONIC SIMULATOR
  A quality control device for the analyzer’s cartridge signal-reading function
COMPONENTS
ANATOMY
OPERATION
CLEANING
i-STAT 1 PRINTER
  Portable printer used to print records from the analyzer

## Slide 11: Anatomy of the i-STAT 1 Analyzer
i-stat 1 learning system | i-stat 1 analyzer system orientation

## Slide 12: Anatomy of the analyzer
Keypad
Display SCREEN
Keypad
Display SCREEN
i-STAT 1 Wireless Analyzer
i-STAT 1 Analyzer
FRONT OF THE ANALYZER
components
ANATOMY
OPERATION
CLEANING

## Slide 13: Anatomy of the analyzer – cont’d
components
ANATOMY
OPERATION
CLEANING
Abc key
Display SCREEN
Scan Key
Number keys
Menu key
Arrow keys
Enter key
Power key
Decimal key
KEYPAD KEY & FUNCTION3 | Description
SCAN | Activates the barcode scanner.
 | Used as a page key to move from one screen to the next and navigate the alphabet when the ABC key is pressed.
ABC | Used to enter alpha characters on data entry screens.
0-9 | Used to enter a number or digits on data entry screen and to select menu options and stored records.
. | Enters a decimal point or a comma separator based on customization.
 | The backlight for the display is turned on and off by pressing the “0” key for one second.
 | Used to respond to a prompt to complete an action, such as entering an operator or patient ID via the keypad.
MENU | Used to return to the previous menu and switch between the Test and Administration Menus.
 | Used to print either directly to the i-STAT 1 Printer or to the i-STAT 1 Printer attached to the i-STAT 1 Downloader/Recharger.
 | Turns the analyzer on or off.
PRINT key
BACKLIGHT key

## Slide 14: Anatomy of the analyzer – cont’d
components
ANATOMY
OPERATION
CLEANING
Abc key
Display SCREEN
Scan Key
Number keys
Menu key
Arrow keys
Enter key
Power key
Decimal key
KEYPAD KEY & FUNCTION3 | Description
SCAN | Activates the barcode scanner.
 | Used as a page key to move from one screen to the next and navigate the alphabet when the ABC key is pressed.
ABC | Used to enter alpha characters on data entry screens.
0-9 | Used to enter a number or digits on data entry screen and to select menu options and stored records.
. | Enters a decimal point or a comma separator based on customization.
 | The backlight for the display is turned on and off by pressing the “0” key for one second.
 | Used to respond to a prompt to complete an action, such as entering an operator or patient ID via the keypad.
MENU | Used to return to the previous menu and switch between the Test and Administration Menus.
 | Used to print either directly to the i-STAT 1 Printer or to the i-STAT 1 Printer attached to the i-STAT 1 Downloader/Recharger.
 | Turns the analyzer on or off.
PRINT key
BACKLIGHT key

## Slide 15: Anatomy of the analyzer – cont’d
Infrared communication window
Laser BARCODE SCAN WINDOW
Battery compartment
TOP OF ANALYZER: i-STAT 1
  Battery compartment
  Infrared communication window
  Laser barcode scan window
components
ANATOMY
OPERATION
CLEANING
CAUTIONS
DO NOT stare into the laser aperture or the laser beam or point the laser beam at others.
Class 2 laser scanners use a low power, visible light diode. As with any bright light source, such as the sun, the user should avoid staring directly into the laser beam. Momentary exposure to a Class 2 laser is not known to be harmful.

## Slide 16: Anatomy of the analyzer – cont’d
Infrared communication window
Laser BARCODE SCAN WINDOW
Battery compartment
TOP OF ANALYZER: i-STAT 1 WIRELESS
  Battery compartment
  Infrared communication window
  Laser barcode scan window
components
ANATOMY
OPERATION
CLEANING
CAUTIONS
DO NOT stare into the laser aperture or the laser beam or point the laser beam at others.
Class 2 laser scanners use a low power, visible light diode. As with any bright light source, such as the sun, the user should avoid staring directly into the laser beam. Momentary exposure to a Class 2 laser is not known to be harmful.

## Slide 17: Anatomy of the analyzer – cont’d
BOTTOM OF ANALYZER: i-STAT 1
  Warning label (i-STAT EC8+ & i-STAT CREA cartridges)
  Cartridge port
WARNING LABEL
CARTRIDGE PORT
CAUTIONS
Propofol, Thiopental Sodium & Hydroxyurea cause interference. See manual for complete interference information.
components
ANATOMY
OPERATION
CLEANING

## Slide 18: Anatomy of the analyzer – cont’d
BOTTOM OF ANALYZER: i-STAT 1 WIRELESS
  Warning label (i-STAT EC8+ & i-STAT CREA cartridges)
  Cartridge port
components
ANATOMY
OPERATION
CLEANING
WARNING LABEL
CARTRIDGE PORT
CAUTIONS
Propofol, Thiopental Sodium & Hydroxyurea cause interference. See manual for complete interference information.

## Slide 19: Anatomy of the analyzer – cont’d
BACK OF ANALYZER
  Model number
  Warning label (barcode scanner)
  Serial number
Model Number
WARNING LABEL
Serial Number
CAUTIONS
DO NOT stare into the laser aperture or the laser beam or point the laser beam at others.
Class 2 laser scanners use a low power, visible light diode. As with any bright light source, such as the sun, the user should avoid staring directly into the laser beam. Momentary exposure to a Class 2 laser is not known to be harmful.
components
ANATOMY
OPERATION
CLEANING

## Slide 20: Batteries | i-STAT 1 Analyzer
BATTERY COMPARTMENT
  Abbott rechargeable battery pack
  9-volt lithium disposable batteries
Abbott rechargeable battery pack
9-volt lithium disposable batteries
Battery Compartment
components
ANATOMY
OPERATION
CLEANING
SCREEN MESSAGES RE: BATTERY
  “Battery Low”
  “Code #1” “Dead Batteries Replace Batteries”

## Slide 21: Batteries | i-STAT 1 Analyzer
BATTERY COMPARTMENT
  Abbott rechargeable battery pack
  9-volt lithium disposable batteries
components
ANATOMY
OPERATION
CLEANING
SCREEN MESSAGES RE: BATTERY
  “Battery Low”
  “Code #1” “Dead Batteries Replace Batteries”
Abbott rechargeable battery pack
9-volt lithium disposable batteries
Battery Compartment

## Slide 22: Basic Operation Process
i-stat 1 learning system | i-stat 1 analyzer system orientation

## Slide 23: Changing the Batteries: 9-volt Disposable | i-STAT 1
FOR 9-VOLT BATTERIES
  Slide open the battery compartment
  Tilt the analyzer to slide out the battery carrier
  Remove the batteries by pushing sideways with thumb
  Insert new batteries into the carrier
  Replace the carrier with gold dots facing down and red dot facing the screen
  Slide battery compartment cover back into place
IMPORTANT NOTE
If the battery cover does not easily slide over the carrier, then the carrier has been installed incorrectly. Remove and ensure proper placement.
components
ANATOMY
OPERATION
CLEANING

## Slide 24: Changing the Batteries: 9-volt Disposable | i-STAT 1 Wireless
FOR 9-VOLT BATTERIES
  Slide open the battery compartment
  Tilt the analyzer to slide out the battery carrier
  Remove the batteries by pushing sideways with thumb
  Insert new batteries into the carrier
  Replace the carrier with gold dots facing down and red dot facing the screen
  Slide battery compartment cover back into place
components
ANATOMY
OPERATION
CLEANING
IMPORTANT NOTE
If the battery cover does not easily slide over the carrier, then the carrier has been installed incorrectly. Remove and ensure proper placement.

## Slide 25: Changing the Batteries: Rechargeable
FOR 9-VOLT BATTERIES
  Slide open the battery compartment
  Tilt the analyzer to slide out the battery pack
  Insert new battery pack with gold dots facing down and the larger red dot facing the screen
  Slide the battery compartment cover back into place
components
ANATOMY
OPERATION
CLEANING
IMPORTANT NOTE
Rechargeable batteries are NOT disposable. The battery pack can be recharged and reused.
Use only a rechargeable battery pack purchased from Abbott.

## Slide 26: Charging the Batteries: Rechargeable – cont’d
RECHARGING THE RECHARGEABLE BATTERY PACK
  In the analyzer
  In the Downloader/Recharger battery compartment
RECHARGING IN THE ANALYZER
  Insert the rechargeable battery pack into the analyzer
  Analyzer must be properly placed in Downloader/Recharger
  Red or green status light illuminates when charging
components
ANATOMY
OPERATION
CLEANING
IMPORTANT NOTE
Batteries can be left in the recharger – they will not overcharge!
Battery compartment
STATUS LIGHT

## Slide 27: Changing the Batteries: Rechargeable – cont’d
RECHARGING IN THE DOWNLOADER/RECHARGER
  Slide the rechargeable battery compartment with:
    Smaller red dot facing up
    Gold contact pads facing the foot
  Green light near the battery compartment illuminates when charging
Battery STATUS LIGHT
components
ANATOMY
OPERATION
CLEANING
IMPORTANT NOTE
Batteries can be left in the recharger – they will not overcharge!

## Slide 28: The Barcode Scanner
USED TO SCAN
  Lot numbers for Quality Tests
  Operator IDs
  Patient IDs
  Cartridge lot number barcode
TO SCAN
  Hold the Analyzer 3 to 9 inches away from the barcode
  Press and hold down the SCAN key
  Laser beam must cover the entire length of the barcode
  Wait for the beep or for laser beam to disappear
  Release the SCAN key
  Repeat if prompted
Barcode Scanner
components
ANATOMY
OPERATION
CLEANING
IMPORTANT NOTE
The barcode scanner is a class II laser device. Momentary exposure is not known to be harmful, however you should never point the scanner at anyone or look into the laser beam.

## Slide 29: The Barcode Scanner – cont’d
IF THE SCAN IS UNSUCCESSFUL, TRY AGAIN
  Make sure the barcode is completely covered by the red beam
  Press and HOLD the SCAN key
  Tilt the barcode away from strong light
components
ANATOMY
OPERATION
CLEANING
IMPORTANT NOTE
The barcode scanner is a class II laser device. Momentary exposure is not known to be harmful, however you should never point the scanner at anyone or look into the laser beam.

## Slide 30: The Infrared Communications Window
INFRARED COMMUNICATION WINDOW
  Designed for communication with:
    Downloaders
    Printers
    Other analyzers
  Accomplished by aligning infrared communication windows
components
ANATOMY
OPERATION
CLEANING
Infrared Communication Window

## Slide 31: Internal Electronic Simulator
INTERNAL ELECTRONIC SIMULATOR
  Performs internal checks to ensure proper and accurate functioning
  Provides independent check on the analyzer’s ability to take accurate and sensitive readings
  Performs test at specified intervals
  Checks whether test needs to be performed each time cartridge is inserted
  Performed automatically when needed before the sample is tested
  Adds 20 seconds to testing cycle
components
ANATOMY
OPERATION
CLEANING

## Slide 32: External Electronic Simulator
EXTERNAL ELECTRONIC SIMULATOR
  The External Electronic Simulator is performed if:
    Required by facility
    Analyzer is dropped or damaged
    There is an Internal Electronic Simulator failure
  Test takes approximately 60 seconds
components
ANATOMY
OPERATION
CLEANING

## Slide 33: Analyzer Display Screen
DISPLAY SCREEN
  Test results
  Entry prompts
  Troubleshooting guidance
components
ANATOMY
OPERATION
CLEANING

## Slide 34: Analyzer Keys: Arrow Keys
ARROW KEYS
  Right Arrow key is for navigating pages within a test result
  Left Arrow key is a backspace to clear keypad entries
components
ANATOMY
OPERATION
CLEANING
Arrow keys

## Slide 35: Analyzer Keys: ABC Key
ABC Key
  Press ABC to enter letter mode
  Use the Arrow keys to scroll through the list
  Select the letter by pressing ABC key again
  If another letter is required, press ABC again to re-enter letter mode
components
ANATOMY
OPERATION
CLEANING
ABC KEY

## Slide 36: Analyzer Keys: Numeric Keys
NUMERIC KEYS
  Used to enter numbers
  Press desired number to enter
components
ANATOMY
OPERATION
CLEANING
IMPORTANT NOTE
Data entry is only possible when prompted.
NUMBER KEYS

## Slide 37: Analyzer Keys: Decimal Point Key
DECIMAL POINT KEYS
  Use . key to enter a decimal point as part of a string of numbers
components
ANATOMY
OPERATION
CLEANING
IMPORTANT NOTE
Data entry is only possible when prompted.
DECIMAL KEY

## Slide 38: Analyzer Keys: “O” Key
“0” Key
  Pressing and holding “O” key will enable backlight
Press & hold “O” for backlight
components
ANATOMY
OPERATION
CLEANING

## Slide 39: Analyzer Keys: Menu Key
MENU KEY
  Toggles between the Test and Administration menus
  Acts as an escape key from any screen back to the menu screen
components
ANATOMY
OPERATION
CLEANING

## Slide 40: Analyzer Cartridge Port – i-STAT 1 Analyzer
CARTRIDGE PORT
  i-STAT cartridges
  Electronic Simulator
components
ANATOMY
OPERATION
CLEANING
Cartridge port

## Slide 41: Analyzer Cartridge Port – i-STAT 1 Wireless Analyzer
CARTRIDGE PORT
  i-STAT cartridges
  Electronic Simulator
components
ANATOMY
OPERATION
CLEANING
Cartridge port

## Slide 42: Test Menu: i-STAT Cartridge
CARTRIDGE TEST MENU OPTIONS
  1 = Last Result
  2 = i-STAT Cartridge (for a patient test)
i-STAT CARTRIDGE
  Press 2 to perform a Patient Test
  Follow the on-screen prompts
components
ANATOMY
OPERATION
CLEANING
IMPORTANT NOTE
Test Menu is used for patient tests only.

## Slide 43: Administration Menu: Analyzer Status
ADMINISTRATION MENU
  Analyzer Status
  Data Review
ANALYZER STATUS SCREEN
  Press MENU to get to the Administration Menu
  Press 1 for Analyzer Status
  Information includes:
    Total records
    Unsent records
    Analyzer Stores 1,000 test records
components
ANATOMY
OPERATION
CLEANING

## Slide 44: Administration Menu: Data Review
DATA REVIEW
  Press Menu to get to the Administration menu
  Press 2 for Data Review
  7 options for grouping and filtering stored data
  View or print any stored record
components
ANATOMY
OPERATION
CLEANING

## Slide 45: Cleaning & Disinfecting
i-stat 1 learning system | i-stat 1 analyzer system orientation

## Slide 46: Cleaning and Disinfecting
PRECAUTIONS
  Always exercise universal safety precautions, when handling the analyzer, cartridges, and peripherals to prevent exposure to blood-born pathogens.
  The use of any unapproved product to clean the i-STAT System may result in damage to system components.
  The analyzer and downloader/recharger are NOT designed to be autoclaved or sterilized by any other method, including high heat, irradiation, or gaseous chemical processes.
COMPONENTS
ANATOMY
OPERATION
CLEANING

## Slide 47: Cleaning and Disinfecting – cont’d
CASE
Display SCREEN
CLEANING THE ANALYZER
  Clean the display screen and the case with any of the following:
    gauze pad moistened with
      isopropyl alcohol (IPA) or
      10% bleach solution
    PDI® Super Sani-Cloth®
  Rinse using a gauze pad moistened with water and dry
CAUTIONS
Avoid getting excess fluids in the seam between the display screen and the case.
The analyzer and downloader/recharger MUST NOT be immersed in any liquid.
COMPONENTS
ANATOMY
OPERATION
CLEANING

## Slide 48: Cleaning and Disinfecting – cont’d
  DECONTAMINATING THE ANALYZER
  Prepare a 10% solution of bleach
  Soak a few gauze pads in the bleach solution. Squeeze the pads to remove excess solution
  Soften, then remove any dried blood with one or two of the gauze pads soaked in the bleach solution. Avoid scraping dried blood as contaminated particles may become airborne
  Clean the entire surface of the device twice with gauze pads soaked in the bleach solution.
CAUTIONS
Decontaminate the analyzer whenever a specimen is spilled onto it or if the item is to be returned to Abbott.
Wear gloves while performing the procedure.
To protect from nosocomial infections, decontaminate analyzers periodically and whenever blood is spilled or transferred to an analyzer.
Avoid getting excess fluids in the seam between the display screen and the case.
COMPONENTS
ANATOMY
OPERATION
CLEANING
  Rinse the surface of the device with gauze pads moistened with tap water and dry. If the device is to be shipped, place it in a plastic bag.

## Slide 49: Cleaning and Disinfecting – cont’d
PLEASE TAKE EXTRA CARE
  Do not let liquid enter the cartridge port or battery compartment; do not immerse in any liquid.
COMPONENTS
ANATOMY
OPERATION
CLEANING
CAUTIONS
Unplug the power supply from the outlet and dry the Downloader completely.

## Slide 50
Congratulations!
You have completed the lesson i-STAT 1 System Orientation