## Slide 1: Please delete this slide after reading through and before working on your slides
TIPS
Content provided in this lesson does not replace the i-STAT 1 User Guides, System Manual, Cartridge Test and Information (CTI) sheets or Instructions for Use (IFU).
This slide deck may be modified to represent products within the i-STAT 1 System currently in use within your facility.
Once the deck has been modified, customers will assume responsibility for the content.*
Abbott will continue to provide and support the original resources that are made available on the www.globalpointofcare.abbott website.*
Not all products are available in all regions.
*Check the website regularly to ensure that you have the most up-to-date content.

## Slide 2
Cartridge Handling
& Sample Collection
i-STAT LEARNING SYSTEM | i-STAT 1 SYSTEM

## Slide 3: Learning objectives
i - STAT LEARNING SYSTEM | GENERAL CARTRIDGE HANDLING & SAMPLE COLLECTION
Describe the basic i-STAT cartridge handling
Describe the proper sample collection and handling practices
List factors that may affect the sample prior to analysis
UNDERSTAND THE FOLLOWING:
SAMPLE COLLECTION
AFFECTING RESULTS
CARTRIDGE OVERVIEW

## Slide 4
i-STAT LEARNING SYSTEM | i-STAT 1 ANALYZER
Cartridge Handling Overview

## Slide 5: i-STAT cartridges
Self-contained lab analyzers
Analysis, sample manipulation, and waste are confined within the cartridge
Unit-use disposable cartridge

## Slide 6: Cartridge box labeling
A Refrigerated storage temperature
indicator: 2-8° C (35-46° F)
B Indicates shelf life when stored
at room temperature
C Refrigerated storage expiration date
D Cartridge LOT number
E Location to record room temperature
expiration date
F Cartridge List Number
A
B
C
D
E
F

## Slide 7: Cartridge pouch labeling
A Cartridge name
B Analyte(s) measured & calculated,
if applicable
C Area to record room temperature expiration date
D 2D barcode for manufacturing quality control; not scannable
E Cartridge LOT number
F Cartridge pouch barcode
G Refrigerated storage expiration date
H Indicates shelf life when stored at room temperature
I Room temperature storage range
FRONT
BACK

## Slide 8: Cartridge portion pack labeling
A Cartridge name
B Analyte(s) measured & calculated,
if applicable
C 2D barcode for manufacturing quality control; not scannable
D Cartridge LOT number
E Cartridge portion pack barcode
F Refrigerated storage expiration date
G Refrigerated temperature storage range
FRONT
BACK
B
C
D
E
F
G
XXX
A

## Slide 9: i-STAT cartridge components
SLIDE CLOSURE
CONTACT PADS & SENSORS
FILL MARK
SAMPLE WELL
CALIBRANT PACK OR ANALYSIS FLUID
WHITE CARTRIDGE | SLIDE CLOSURE
SLIDE-CLOSURE CARTRIDGES
DO NOT TOUCH

## Slide 10: i-STAT cartridge components
WHITE CARTRIDGE | SNAP CLOSURE
CONTACT PADS & SENSORS
FILL MARK
CALIBRANT PACK
SAMPLE WELL
SNAP CLOSURE
DO NOT TOUCH
SNAP-CLOSURE CARTRIDGES

## Slide 11: i-STAT cartridge components:
APPLICABLE CARTRIDGES
WHITE COAGULATION CARTRIDGE | SNAP CLOSURE
CONTACT PADS & SENSORS
DO NOT TOUCH
SNAP CLOSURE
SAMPLE WELL
FILL MARK

## Slide 12: i-STAT cartridge components:
BLUE CARTRIDGES
SNAP CLOSURE
CONTACT PADS & SENSORS
FILL MARK
CALIBRANT PACK
DO NOT TOUCH
SAMPLE WELL
BLUE CARTRIDGE | SNAP CLOSURE

## Slide 13: Storage requirements
REFRIGERATED STORAGE
May be used up to expiration date
Refrigerated storage is 2°C - 8°C
ROOM TEMPERATURE STORAGE
14d = 14 days
2m = 2 months
Room temperature storage range is 18°C - 30°C
EXAMPLE CARTRIDGES

## Slide 14: Preparation for use
CARTRIDGES MUST BE BROUGHT TO ROOM TEMPERATURE BEFORE USE
5 minutes for a single cartridge
60 minutes for a box of cartridges
CAUTIONS Do NOT attempt to speed up the cartridge’s transition to room temperature

## Slide 15: Preparation for use
AFTER REMOVING CARTRIDGE(S) FROM REFRIGERATION
Must not be returned to the refrigerator.
Record the room temperature expiration date on the box or individual cartridge.
If the box is labeled, keep cartridges in the box until use.
All cartridges should be used immediately after opening pouch.
If the pouch has been punctured, the cartridge should not be used
Although the sample is contained within the cartridge, cartridges should be disposed of as biohazardous waste according to local, state, and national regulatory guidelines.
RECORD THE ROOM TEMPERATURE EXPIRATION DATE ON BOX & POUCH HERE
DISPOSE AFTER USE:
PLACE THE CARTRIDGE IN A DESIGNATED BIOHAZARD WASTE CONTAINER

## Slide 16: Handling white cartridges - pouches
Carefully tear open the cartridge pouch where indicated
Hold the cartridge by the sides or the bottom when filling and handling
IMPORTANT: To prevent a Quality Check Code, avoid touching the sensors or pressing the center of the cartridge.

## Slide 17: Handling white cartridges – portion pack
Carefully tear open the portion pack where indicated by the arrows
Hold the cartridge by the sides or the bottom when filling and handling
IMPORTANT: To prevent a Quality Check Code, avoid touching the sensors or pressing the center of the cartridge.
FRONT
BACK
TEAR AT ARROWS

## Slide 18: Handling blue cartridges - pouches
Carefully tear open the cartridge pouch where indicated
Hold the cartridge by the sides or the bottom when filling/handling
IMPORTANT: To prevent a Quality Check Code, avoid touching the sensors or pressing the center of the cartridge.

## Slide 19
i-STAT LEARNING SYSTEM | i-STAT 1 ANALYZER
General Sample Collection Techniques

## Slide 20: Sample Types
i-STAT SYSTEM CAN BE USED WITH THE FOLLOWING SAMPLE TYPES:
Arterial | Venous | Capillary | Plasma
REFER TO THE INTENDED USE STATEMENTS FOR EACH i-STAT PRODUCT:
Instructions for Use (IFU) and Cartridge and Test Information (CTI) Sheets can be found in the "i-STAT Support" area of the www.globalpointofcare.abbott website.
ARTERIAL SAMPLES
VENOUS SAMPLES
CAPILLARY SAMPLES
PLASMA SAMPLES
IMPORTANT: Follow your hospitals procedures for proper sample collection. Correlate the results with the patient’s status.

## Slide 21: Sample collection | What to avoid
  IV LINE
    Avoid drawing from an arm with an IV line
      - IV solutions dilute sample
  INDWELLING LINE
    Draw sufficient waste
    Do not introduce air into the sample
  TOURNIQUET
    If applied for more than 1 minute, release and reapply after 2 to 3 minutes
AVOID EXTRA MUSCULAR ACTIVITY:
No clenching and unclenching of fist
AVOID HEMOLYSIS:
Allow residual alcohol to dry
Discard sample from traumatic draw
ALLOW TO DRY

## Slide 22: Sample collection | Capillary
CAPILLARY TUBE
Use a skin puncture device that provides free-flowing blood
Use a balanced heparin capillary tube
SKIN PUNCTURE DEVICE
EXAMPLE OF HEPARINIZED CAPILLARY TUBE

## Slide 23
WHEN PERFORMING A FINGERSTICK
Use a skin puncture device that provides free-flowing blood
Gently squeeze the finger to produce a hanging drop of blood
Sample collection | Fingerstick
SKIN PUNCTURE DEVICE
IMPORTANT: Direct fingerstick application is only recommended for the i-STAT PTplus cartridge

## Slide 24
i-STAT LEARNING SYSTEM | i-STAT 1 ANALYZER
Factors that May Affect Results

## Slide 25: Avoiding factors that may affect the sample prior to analysis
ALWAYS CORRELATE RESULTS WITH THE STATUS OF THE PATIENT
REFER TO:
  Instructions for Use (IFU) and Cartridge and Test Information Sheets (CTI Sheets)
  i‑STAT 1 System Manual
  www.globalpointofcare.abbott

## Slide 26: Factors affecting Sodium results
AVOID AFFECTING SODIUM RESULTS
  Draw adequate waste prior to drawing blood
  Avoid drawing blood from an arm with an IV line
IMPORTANT: Mix thoroughly and test promptly

## Slide 27: Factors affecting Potassium results
AVOID AFFECTING POTASSIUM RESULTS
  Prevent hemolysis resulting from traumatic draws / vigorous sample handling
  Do NOT ice samples
  Test promptly
  Avoid extra muscular activity when drawing blood
HEMOLYSIS
TEST PROMPTLY
IMPORTANT: Mix thoroughly and test promptly

## Slide 28: Factors affecting Ionized Calcium results
AVOID AFFECTING IONIZED CA RESULTS
  Fill blood collection devices completely
  Limit extra muscular activity during sample collection
IMPORTANT: Mix thoroughly and test promptly

## Slide 29: Factors affecting Glucose results
AVOID AFFECTING GLUCOSE RESULTS
  Test Promptly
  If the patient has been administered hydroxyurea (Droxia®, Hydrea®), note increased results ≥ 0.08 mmol/L
  Bromide (Lithium Bromide) at ≥ 11.8 mmol/L showed decreased results
HYDROXYUREA
TEST PROMPTLY
IMPORTANT: Mix thoroughly and test promptly
BROMIDE
(Lithium Bromide)

## Slide 30: Factors affecting Creatinine results
AVOID AFFECTING CREATININE RESULTS
  Use another testing method if the patient has been administered hydroxyurea (Droxia®, Hydrea®)
HYDROXYUREA
IMPORTANT: Mix thoroughly and test promptly

## Slide 31: Factors affecting Lactate results
AVOID AFFECTING LACTATE RESULTS
  Test immediately
  Do NOT use a tourniquet
  Use another testing method if the patient has been administered hydroxyurea (Droxia®, Hydrea®) OR Bromide (Lithium Bromide)
TEST PROMPTLY
HYDROXYUREA
BROMIDE
(Lithium Bromide)
IMPORTANT: Mix thoroughly and test promptly

## Slide 32: Factors affecting Hematocrit results
AVOID AFFECTING HEMATOCRIT RESULTS
  Mix thoroughly
  Draw adequate waste prior to drawing blood
  Keep analyzer level during testing
IMPORTANT: Mix thoroughly and test promptly
5x
10x

## Slide 33: Factors affecting pH results
AVOID AFFECTING PH RESULTS
  Avoid extra muscular activity during sample collection
  Do NOT expose to air
IMPORTANT: Mix thoroughly and test promptly

## Slide 34: Factors affecting PCO2 results in ABG cartridges
AVOID AFFECTING PCO2 RESULTS
  Test promptly
  Fill blood collection device completely
  Do NOT expose to air
TEST PROMPTLY
IMPORTANT: Mix thoroughly and test promptly

## Slide 35: Factors affecting PO2 results in ABG cartridges
AVOID AFFECTING PO2 RESULTS
  Test promptly
  Do NOT ice samples
  Fill blood collection device completely
  Do NOT expose to air
TEST PROMPTLY
IMPORTANT: Mix thoroughly and test promptly

## Slide 36: Factors affecting ACT results
AVOID AFFECTING ACT RESULTS
Do NOT use glass tubes
Ensure blood collection device does not contain anticoagulants, clot activators, or serum separators
Level testing surface
Flush lines sufficiently
IMPORTANT: Mix thoroughly and test promptly
GLASS TUBES

## Slide 37: Factors affecting PT results
AVOID AFFECTING RESULTS
Use another method if the patient is receiving CUBICIN® (daptomycin for injection)
Place the handheld on a level, nonvibrating surface
Flush lines sufficiently
Ensure blood collection device does not contain anticoagulants, clot activators or separators
IMPORTANT: Mix thoroughly and test promptly
CUBICIN®
(daptomycin for injection)

## Slide 38: Factors affecting Troponin results
AVOID AFFECTING TROPONIN RESULTS
Keep analyzer on a level surface during the testing process
Do NOT move the analyzer during the testing process
Mix thoroughly
Prevent hemolysis
IMPORTANT: Mix thoroughly and test promptly
HEMOLYSIS
5x
10x

## Slide 39: Summary
NOW THAT YOU HAVE COMPLETED THIS LESSON, YOU SHOULD BE ABLE TO:
  Describe the basic i-STAT cartridge-handling guidelines
  Describe proper sample collection and handling practices
  List factors that may affect the sample prior to analysis
i - STAT LEARNING SYSTEM | GENERAL CARTRIDGE HANDLING & SAMPLE COLLECTION

## Slide 40: Summary
Congratulations!
Congratulations!
You have completed the lesson General Cartridge Handling & Sample Collection