## Slide 1: Please delete this slide after reading through and before working on your slides
TIPS
Content provided in this lesson does not replace the i-STAT 1 User Guide, i-STAT 1 System Manual or Instructions for Use (IFU).
This power point may be modified to represent products within the i-STAT 1 System currently in use within your facility.
Once the resource has been modified, customers will then assume responsibility for the content.*
Abbott will continue to provide and support the original resources made available on the www.globalpointofcare.abbott website.*
Not all products are available in all regions.
*Check the website regularly to ensure that you have the most up-to-date content.

## Slide 2: i-STAT ACTc / ACTk Cartridges
PERFORMING A PATIENT TEST | i-STAT 1 SYSTEM

## Slide 3: Learning objectives
PERFORMING A PATIENT TEST | i-STAT ACTc / ACTk CARTRIDGE
UNDERSTAND THE FOLLOWING:
CARTRIDGE OVERVIEW
RUNNING A TEST
AVOIDING ERRORS
Describe the basic i-STAT cartridge handling
Describe the steps of running an i-STAT test
Troubleshooting and factors affecting samples prior to analysis

## Slide 4
PERFORMING A PATIENT TEST
i-STAT ACTc / ACTk Cartridges

## Slide 5: ACTc Cartridge overview

## Slide 6: ACTc Cartridge overview– cont’d
INTENDED USE
The i-STAT Celite Activated Clotting Time (CeliteACT) test is an in vitro diagnostic test that uses fresh, whole blood, and is useful for monitoring patients receiving heparin for treatment of pulmonary embolism or venous thrombosis, and for monitoring anticoagulation therapy in patients undergoing medical procedures such as catheterization, cardiac surgery, surgery, organ transplant, and dialysis.
SYSTEM COMPATIBILITY
The i-STAT Celite ACT cartridge is compatible with the i-STAT 1 System

## Slide 7: ACTc Cartridge overview– cont’d
CARTRIDGE INFORMATION ON POUCH
  On each cartridge packaging, there is important information to locate:
  A Cartridge name
  B Analytes – measured and calculated, if applicable
  C Location to record room temperature expiration date
  D 2D barcode for manufacturing quality control; not scannable
  E Cartridge LOT number
  F Cartridge pouch barcode
  G Refrigerated storage expiration date
  H Indicates shelf life when stored at room temperature
  I Room temperature storage range
FRONT
BACK
A
B
C
D
E
F
G
H
I

## Slide 8: ACTc Cartridge overview– cont’d
ANATOMY OF A CARTRIDGE
On each cartridge, there are important areas to note
  A Contact pads & sensors (do not touch)
  B Calibrant Pack (do not touch)
  C Cartridge Closure
  D Fill To Mark
  E Sample Well
  F Sample Chamber
A
B
C
F
D
E
TOP
BOTTOM

## Slide 9: ACTk Cartridge overview

## Slide 10: ACTk Cartridge overview– cont’d
INTENDED USE
The i-STAT Kaolin Activated Clotting Time (KaolinACT) test is an in vitro diagnostic test that uses fresh, whole blood, and is used to monitor high-dose heparin anticoagulation frequently associated with cardiovascular surgery.
SYSTEM COMPATIBILITY
The i-STAT Kaolin ACT cartridge is compatible with the i-STAT 1 System

## Slide 11: ACTk Cartridge overview– cont’d
CARTRIDGE INFORMATION ON POUCH
  On each cartridge packaging, there is important information to locate:
  A Cartridge name
  B Analytes – measured and calculated, if applicable
  C Location to record room temperature expiration date
  D 2D barcode for manufacturing quality control, not scannable
  E Cartridge LOT number
  F Cartridge pouch barcode
  G Refrigerated storage expiration date
  H Indicates shelf life when stored at room temperature
  I Room temperature storage range
FRONT
BACK
A
B
C
D
E
F
G
H
I

## Slide 12: ACTk Cartridge overview– cont’d
A
B
C
F
D
E
TOP
BOTTOM
ANATOMY OF A CARTRIDGE
On each cartridge, there are important areas to note
  A Contact pads & sensors (do not touch)
  B Calibrant Pack (do not touch)
  C Cartridge Closure
  D Fill To Mark
  E Sample Well
  F Sample Chamber

## Slide 13
PERFORMING A PATIENT TEST
Running a Test

## Slide 14: Before beginning
QUICK TIPS
For best results, the cartridge and analyzer should be at room temperature
A cartridge should not be removed from its protective pouch until it is at room temperature (18-30 °C or 64-86 °F)
Since condensation on a cold cartridge may prevent proper contact with the analyzer, allow refrigerated cartridges to equilibrate at room temperature for 5 minutes for a single cartridge and 1 hour for an entire box before use

## Slide 15: Before beginning – cont’d
PRECAUTIONS
Use a cartridge immediately after removing it from its protective pouch. Prolonged exposure may cause a cartridge to fail a Quality Check
Do not return unopened, previously refrigerated cartridges to the refrigerator
Cartridges may be stored at room temperature for the time frame indicated on the cartridge box

## Slide 16: Testing process overview
| 16
FLUSH LINE
COLLECT SAMPLE
FILL & CLOSE CARTRIDGE
INSERT CARTRIDGE
3
2
1
5
4
PREPARE ANALYZER

## Slide 17: Prepare the analyzer – i-STAT 1
POWER UP THE ANALYZER
Press the power button to turn on the analyzer
From the Test Menu screen, press 2 on the keypad for i-STAT Cartridge
Follow the on-screen prompts
| 17
PRESS THE POWER BUTTON TO TURN ON THE i-STAT 1
PRESS TO PERFORM A PATIENT TEST

## Slide 18: Prepare the analyzer – i-STAT 1 Wireless
POWER UP THE ANALYZER
Press the power button to turn on the analyzer
From the Test Menu screen, press 2 on the keypad for i-STAT Cartridge
Follow the on-screen prompts
| 18
PRESS THE POWER BUTTON TO TURN ON THE i-STAT 1
PRESS TO PERFORM A PATIENT TEST

## Slide 19: Prepare the analyzer – cont’d
BARCODE SCAN: OPERATOR ID, THEN PATIENT ID
Hold the analyzer 3 to 9 inches away from the barcode
Press and hold down the SCAN key
Laser beam must cover the entire length of the barcode
Wait for the beep or for laser beam to disappear
Release the SCAN key
Repeat if prompted
OR ENTER THE PATIENT ID
or enter your Operator ID
OR ENTER THE OPERATOR ID
IMPORTANT
DO NOT stare into the laser aperture or the laser beam or point the laser beam at others.
Class 2 laser scanners use a low power, visible light diode. As with any bright light source, such as the sun, the user should avoid staring directly into the laser beam. Momentary exposure to a Class 2 laser is not known to be harmful.

## Slide 20: Prepare the analyzer – i-STAT 1
PATIENT ID
OPERATOR ID
MANUAL ENTRY: OPERATOR ID, THEN PATIENT ID
Use number and/or ABC keys on the keypad​
  Up to 15 alphanumeric characters may be entered
Press ENT key
Repeat if prompted

## Slide 21: Prepare the analyzer – i-STAT 1 Wireless
MANUAL ENTRY: OPERATOR ID, THEN PATIENT ID
Use number and/or ABC keys on the keypad​
  Up to 15 alphanumeric characters may be entered
Press ENT key
Repeat if prompted
OPERATOR ID
PATIENT ID

## Slide 22: Prepare the analyzer – cont’d
SCAN CARTRIDGE (POUCH) BARCODE
Once the cartridge barcode is scanned, you have 15 minutes to insert a filled & sealed cartridge
After 15 minutes of inactivity the analyzer will turn off and will require re-entry of previous information to perform a test
Follow the handheld prompts to prepare the cartridge and perform testing
PRECAUTIONS
If an Invalid Cartridge Type message is displayed, contact the System Administrator.
Scanning is required. You must scan the barcode. This information cannot be entered manually.
or enter your Operator ID
or enter your Operator ID
THE BARCODE ON THE CARTRIDGE POUCH

## Slide 23: Prepare the analyzer – cont’d
SCAN (CARTRIDGE POUCH) BARCODE – ACTc
Scan the barcode on the cartridge pouch
The lot information displays on the analyzer screen momentarily
The analyzer is now ready, next prepare the same
hs-TnI
FRONT
BACK
PRECAUTIONS
If an Invalid Cartridge Type message is displayed, contact the System Administrator.
Scanning is required. You must scan the barcode. This information cannot be entered manually.
SCAN THE LINEAR BARCODE​
DO NOT SCAN THE 2D BARCODE​

## Slide 24: Prepare the analyzer – cont’d
SCAN (CARTRIDGE POUCH) BARCODE – ACTk
Scan the barcode on the cartridge pouch
The lot information displays on the analyzer screen momentarily
The analyzer is now ready, next prepare the same
hs-TnI
FRONT
BACK
SCAN THE LINEAR BARCODE​
DO NOT SCAN THE 2D BARCODE​
PRECAUTIONS
If an Invalid Cartridge Type message is displayed, contact the System Administrator.
Scanning is required. You must scan the barcode. This information cannot be entered manually.

## Slide 25: Prepare for sample collection
QUICK TIPS
The i-STAT ACTc and i-STAT ACTk cartridges require fresh, venous or arterial whole blood (approximately 40μL)
DO
DO use a collection technique resulting in good blood flow
DO ensure the sample for testing is drawn into a plastic collection device (either a plastic syringe or plastic evacuated tube).
DO fill the cartridge immediately to the ‘fill’ line. Delay in filling may produce erroneous results
DO NOT
DO NOT use any collection device that contains anticoagulants
DO NOT use any collection device that has clot activators or serum separators
DO NOT incorrectly handle or incorrectly fill the cartridge

## Slide 26: Prepare for sample collection
BLOOD COLLECTION OPTIONS
ACT ANALYTE |  | SYRINGES | TEST TIMING |  | EVACUATED TUBES | TEST TIMING | 
ACTc, ACTk |  | Plastic syringe without anticoagulant, clot activators, or serum separators | Immediately |  | Plastic collection tube without anticoagulant, clot activators, or serum separators | Immediately | 

## Slide 27: Running the test
REMOVE THE CARTRIDGE FROM THE CARTRIDGE POUCH - ACTc
Place the cartridge on a clean, flat, level, and vibration-free surface
Handle the cartridge by the sides
PRECAUTIONS
DO NOT touch the sensors on the top of the cartridge.
DO NOT apply pressure to the central area of the label because the calibration fluid pack inside could burst prematurely.
DO NOT expose cartridges to temperatures above 30°C (86°F).
DO NOT place the cartridge and/or analyzer in an oxygen-enriched atmosphere.
DO NOT
TOUCH THE SENSORS
DO NOT TOUCH THE CENTER OF THE CARTRIDGE

## Slide 28: Running the test
REMOVE THE CARTRIDGE FROM THE CARTRIDGE POUCH - ACTk
Place the cartridge on a clean, flat, level, and vibration-free surface
Handle the cartridge by the sides
PRECAUTIONS
DO NOT touch the sensors on the top of the cartridge.
DO NOT apply pressure to the central area of the label because the calibration fluid pack inside could burst prematurely.
DO NOT expose cartridges to temperatures above 30°C (86°F).
DO NOT place the cartridge and/or analyzer in an oxygen-enriched atmosphere.
DO NOT
TOUCH THE SENSORS
DO NOT TOUCH THE CENTER OF THE CARTRIDGE

## Slide 29: Running the test – cont’d
FRESH WHOLE BLOOD: INDWELLING LINE
Fluid drip through the line must be discontinued. The line should be flushed with 5 mL of saline and the first 5 mL of blood or six dead space volumes should be discarded
Withdraw the sample for testing into a fresh plastic syringe
The collection syringe cannot contain anticoagulants such as heparin, EDTA, oxalate, or citrate
Dispense sample immediately into the sample well of a cartridge
PRECAUTIONS
Cautions should be used when handling materials that may contain transmissible infectious agents.
Fill evacuated tube per manufacturer’s recommendations.

## Slide 30: Running the test – cont’d
FRESH WHOLE BLOOD: EXTRACORPOREAL LINE
Flush the extracorporeal blood access line by withdrawing 5 mL of blood into a syringe and discard the syringe
Withdraw the sample for testing into a fresh plastic syringe
The collection syringe cannot contain anticoagulants such as heparin, EDTA, oxalate, or citrate
Dispense sample immediately into the sample well of a cartridge
PRECAUTIONS
Cautions should be used when handling materials that may contain transmissible infectious agents.
Fill evacuated tube per manufacturer’s recommendations.

## Slide 31: Running the test – cont’d
WHOLE BLOOD: FILL THE i-STAT CARTRIDGE
Fill the cartridge immediately after collection
Direct the hub of syringe or tip of the transfer device (pipette or dispensing tip) into the sample well of the cartridge
Slowly dispense sample into the sample well until the sample reaches the fill mark indicated on the cartridge. Cartridge is properly filled when the sample reaches the ‘fill to’ mark and a small amount of sample is in the sample well.
CORRECTLY FILLED TO FILL LINE

## Slide 32: Running the test – cont’d
CLOSE AND SEAL THE i-STAT CARTRIDGE
Fold the snap closure over the sample well to close and seal the cartridge
PRECAUTIONS
Incorrect handling and filling of the i-STAT cartridge will result in a quality check failure.
    CORRECTLY FILLED CARTRIDGE
    UNDER-FILLED CARTRIDGE
    OVER-FILLED CARTRIDGE

## Slide 33: Running the test – cont’d
INSERT THE CARTRIDGE
Insert the cartridge into the i-STAT analyzer’s cartridge port until it clicks into place.
  Sensors facing up, and on leading edge
PRECAUTIONS
DO NOT touch the sensors on the top of the cartridge.
DO NOT apply pressure to the central area of the label because the calibration fluid pack inside could burst prematurely.
DO NOT place the cartridge and/or analyzer in an oxygen-enriched atmosphere.

## Slide 34: Running the test – cont’d
PRIOR TO TESTING
The analyzer will perform quality checks
Failure of a quality check will result in a Quality Check Code

## Slide 35: Running the test – cont’d
TIME TO RESULTS
Once the cartridge is inserted, Contacting Cartridge will display followed by the countdown bar
The i-STAT ACTc & i-STAT ACTk cartridges’ time to detection of end point - up to 1000 sec (16.7 min)
This allows the user to estimate the time to results
Alerts, such as “Cartridge Locked,” are also displayed
PRECAUTIONS
DO NOT attempt to remove the cartridge while the message “Cartridge Locked” remains on the screen.
ACT-C
Test In
Progress
▪ ▪ ▪
>260 Sec
1 – Stop Test

## Slide 36: Running the test – cont’d
TIME TO RESULTS
Once the cartridge is inserted, Contacting Cartridge will display followed by the countdown bar
The i-STAT ACTc & i-STAT ACTk cartridges’ time to detection of end point - up to 1000 sec (16.7 min)
This allows the user to estimate the time to results
Alerts, such as “Cartridge Locked,” are also displayed
PRECAUTIONS
DO NOT attempt to remove the cartridge while the message “Cartridge Locked” remains on the screen.
ACT-C
Test In
Progress
▪ ▪ ▪
>260 Sec
1 – Stop Test

## Slide 37
PERFORMING A PATIENT TEST
Avoiding Errors

## Slide 38: Troubleshooting and avoiding errors
PREVENT ERROR MESSAGES AND QUALITY CHECK CODES
There are potential pre-analytical conditions, especially during the training period, that may not be detected by the operator.
  The analyzer will detect the condition, halt the test cycle, and display a cause message followed by the action message “USE ANOTHER CARTRIDGE.”
The cartridge and analyzer must be at room temperature (18-30 °C or 64-86 °F) to perform the test.
The cartridge must be correctly filled, closed and sealed.
The analyzer must remain on a level surface. Level surface includes performing the test with the analyzer in the downloader/recharger.

## Slide 39: Factors affecting ACT results
AVOID AFFECTING ACT RESULTS
Do NOT use glass tubes
Ensure blood collection device does not contain anticoagulants, clot activators, or serum separators
Level testing surface
Flush lines sufficiently
IMPORTANT
Mix thoroughly and test promptly
GLASS TUBES

## Slide 40
Congratulations!
You have completed the lesson
i-STAT 1 ACTc & ACTk
Cartridges