## Slide 1: Please delete this slide after reading through and before working on your slides
TIPS
Content provided in this lesson does not replace the i-STAT 1 User Guide, i-STAT 1 System Manual or Instructions for Use (IFU).
This power point may be modified to represent products within the i-STAT 1 System currently in use within your facility.
Once the resource has been modified, customers will then assume responsibility for the content.*
Abbott will continue to provide and support the original resources made available on the www.globalpointofcare.abbott website.*
Not all products are available in all regions.
*Check the website regularly to ensure that you have the most up-to-date content.

## Slide 2: i-STAT EG6+/ EG7+ Cartridge USA Only
PERFORMING A PATIENT TEST | i-STAT 1 SYSTEM

## Slide 3: Learning objectives
PERFORMING A PATIENT TEST | i-STAT EG6+ / EG7+ CARTRIDGE USA ONLY
Describe the basic i-STAT cartridge handling
Describe the steps of running an i-STAT test
Troubleshooting and factors affecting samples prior to analysis
UNDERSTAND THE FOLLOWING:
CARTRIDGE OVERVIEW
RUNNING A TEST
AVOIDING ERRORS

## Slide 4
PERFORMING A PATIENT TEST
i-STAT EG6+/ EG7+ Cartridge USA Only

## Slide 5: EG6+ Cartridge overview

## Slide 6: EG6+ Cartridge overview - cont’d
INTENDED USE
The i-STAT EG6+ cartridge with the i-STAT 1 System is intended for use in the in vitro quantification of Sodium, Potassium, Hematocrit, pH, oxygen partial pressure, and carbon dioxide partial pressure in arterial, venous or capillary* whole blood.
SYSTEM COMPATIBILITY
The i-STAT EG6+ cartridge is compatible with the i-STAT 1 System
IMPORTANT
*Capillary samples not applicable for Potassium tests.

## Slide 7: EG6+ Cartridge overview - cont’d
CARTRIDGE INFORMATION ON POUCH
  On each cartridge packaging, there is important information to locate:
  A Cartridge name
  B Analytes – measured and calculated, if applicable
  C Location to record room temperature expiration date
  D 2D barcode for manufacturing quality control; not scannable
  E Cartridge LOT number
  F Cartridge pouch barcode
  G Refrigerated storage expiration date
  H Indicates shelf life when stored at room temperature
  I Room Temperature storage range
FRONT
BACK
A
B
C
D
E
F
G
H
I

## Slide 8: EG6+ Cartridge overview - cont’d
ANATOMY OF A CARTRIDGE
On each cartridge, there are important areas to note
  A Contact pads & sensors (do not touch)
  B Calibrant Pack (do not touch)
  C Cartridge Closure
  D Fill To Mark
  E Sample Well
  F Sample Chamber
TOP
BOTTOM
A
B
C
D
E
F

## Slide 9: EG7+ Cartridge overview

## Slide 10: EG7+ Cartridge overview - cont’d
INTENDED USE
The i-STAT EG7+ cartridge with the i-STAT 1 System is intended for use in the in vitro quantification of Sodium, Potassium, Ionized Calcium, Hematocrit, pH, oxygen partial pressure, and carbon dioxide partial pressure in arterial, venous or capillary* whole blood.
SYSTEM COMPATIBILITY
The i-STAT EG7+ cartridge is compatible with the i-STAT 1 System
IMPORTANT
*Capillary samples not applicable for Potassium and Ionized Calcium tests.

## Slide 11: EG7+ Cartridge overview - cont’d
FRONT
BACK
A
B
C
D
E
F
G
H
I
CARTRIDGE INFORMATION ON POUCH
  On each cartridge packaging, there is important information to locate:
  A Cartridge name
  B Analytes – measured and calculated, if applicable
  C Location to record room temperature expiration date
  D 2D barcode for manufacturing quality control; not scannable
  E Cartridge LOT number
  F Cartridge pouch barcode
  G Refrigerated storage expiration date
  H Indicates shelf life when stored at room temperature
  I Room temperature storage range

## Slide 12: EG7+ Cartridge overview - cont’d
TOP
BOTTOM
A
B
C
D
E
F
ANATOMY OF A CARTRIDGE
On each cartridge, there are important areas to note
  A Contact pads & sensors (do not touch)
  B Calibrant Pack (do not touch)
  C Cartridge Closure
  D Fill To Mark
  E Sample Well
  F Sample Chamber

## Slide 13
PERFORMING A PATIENT TEST
Running a Test

## Slide 14: Before beginning
QUICK TIPS
For best results, the cartridge and analyzer should be at room temperature
A cartridge should not be removed from its protective pouch until it is at room temperature (18-30 °C or 64-86 °F)
Since condensation on a cold cartridge may prevent proper contact with the analyzer, allow refrigerated cartridges to equilibrate at room temperature for 5 minutes for a single cartridge and 1 hour for an entire box before use

## Slide 15: Before beginning – cont’d
PRECAUTIONS
Use a cartridge immediately after removing it from its protective pouch. Prolonged exposure may cause a cartridge to fail a Quality Check
Do not return unopened, previously refrigerated cartridges to the refrigerator
Cartridges may be stored at room temperature for the time frame indicated on the cartridge box

## Slide 16: Testing process overview
| 16
COLLECT SAMPLE
MIX SAMPLE
DISCARD FIRST DROPS
FILL & CLOSE CARTRIDGE
INSERT CARTRIDGE
PREPARE ANALYZER
10x
5x
1
2
3
4
5
6

## Slide 17: Prepare the analyzer - i-STAT 1
POWER UP THE ANALYZER
Press the power button to turn on the analyzer
From the Test Menu screen, press 2 on the keypad for “i-STAT Cartridge”
Follow the on-screen prompts
| 17
PRESS THE POWER BUTTON TO TURN ON THE i-STAT 1
PRESS TO PERFORM A PATIENT TEST

## Slide 18: Prepare the analyzer - i-STAT 1 Wireless
POWER UP THE ANALYZER
Press the power button to turn on the analyzer
From the Test Menu screen, press 2 on the keypad for “i-STAT Cartridge”
Follow the on-screen prompts
| 18
PRESS THE POWER BUTTON TO TURN ON THE i-STAT 1
PRESS TO PERFORM A PATIENT TEST

## Slide 19: Prepare the analyzer – cont’d
BARCODE SCAN: OPERATOR ID, THEN PATIENT ID
Hold the analyzer 3 to 9 inches away from the barcode
Press and hold down the SCAN key
Laser beam must cover the entire length of the barcode
Wait for the beep or for laser beam to disappear
Release the SCAN key
Repeat if prompted
OR ENTER THE PATIENT ID
or enter your Operator ID
OR ENTER THE OPERATOR ID
IMPORTANT
DO NOT stare into the laser aperture or the laser beam or point the laser beam at others.
Class 2 laser scanners use a low power, visible light diode. As with any bright light source, such as the sun, the user should avoid staring directly into the laser beam. Momentary exposure to a Class 2 laser is not known to be harmful.

## Slide 20: Prepare the analyzer – i-STAT 1
PATIENT ID
OPERATOR ID
MANUAL ENTRY: OPERATOR ID, THEN PATIENT ID
Use number and/or ABC keys on the keypad​
  Up to 15 alphanumeric characters may be entered
Press ENT key
Repeat if prompted

## Slide 21: Prepare the analyzer – i-STAT 1 Wireless
OPERATOR ID
PATIENT ID
MANUAL ENTRY: OPERATOR ID, THEN PATIENT ID
Use number and/or ABC keys on the keypad​
  Up to 15 alphanumeric characters may be entered
Press ENT key
Repeat if prompted

## Slide 22: Prepare the analyzer – cont’d
SCAN CARTRIDGE (POUCH) BARCODE
Once the cartridge barcode is scanned, you have 15 minutes to insert a filled & sealed cartridge
After 15 minutes of inactivity, the analyzer will turn off and will require re-entry of previous information to perform a test
Follow the on-screen prompts to prepare the cartridge and perform testing
PRECAUTIONS
If an Invalid Cartridge Type message is displayed, contact the System Administrator.
Scanning is required. You must scan the barcode. This information cannot be entered manually.
or enter your Operator ID
or enter your Operator ID
THE BARCODE ON THE CARTRIDGE POUCH
MINS

## Slide 23: Prepare the analyzer – cont’d
SCAN CARTRIDGE (POUCH) BARCODE – EG6+
Scan the barcode on the cartridge pouch
The lot information displays on the analyzer screen momentarily
The analyzer is now ready, next prepare the sample
PRECAUTIONS
If an Invalid Cartridge Type message is displayed, contact the System Administrator.
Scanning is required. You must scan the barcode. This information cannot be entered manually.
SCAN THE LINEAR BARCODE​
DO NOT SCAN THE 2D BARCODE​
FRONT
BACK

## Slide 24: Prepare the analyzer – cont’d
SCAN CARTRIDGE (POUCH) BARCODE – EG7+
Scan the barcode on the cartridge pouch
The lot information displays on the analyzer screen momentarily
The analyzer is now ready, next prepare the sample
PRECAUTIONS
If an Invalid Cartridge Type message is displayed, contact the System Administrator.
Scanning is required. You must scan the barcode. This information cannot be entered manually.
SCAN THE LINEAR BARCODE​
DO NOT SCAN THE 2D BARCODE​
FRONT
BACK

## Slide 25: Prepare for sample collection
QUICK TIPS
The i-STAT EG6+ & i-STAT EG7+ cartridges require arterial, venous or capillary* whole blood (approximately 95μL)
DO
DO use a collection technique resulting in good blood flow
DO collect a venous specimen ensuring proper order of draw & fill a lithium heparin or nonanticoagulated blood collection tube to capacity
DO fill the cartridge to the ‘fill’ line immediately. Delay in filling may produce erroneous results
DO NOT
DO NOT collect a sample from an arm with an IV
DO NOT underfill a blood collection tube
DO NOT incorrectly handle or incorrectly fill the cartridge
* DO NOT use capillary samples for Ionized Calcium and Potassium tests

## Slide 26: Prepare for sample collection - cont’d
BLOOD COLLECTION OPTIONS (part 1 of 2)
Analyte | SYRINGES | Test Timing |  | EVACUATED TUBES | Test Timing | CAPILLARY* | Test Timing
pH, PCO2, PO2, iCa (on EG7+) | Without anticoagulant | 3 minutes |  | Without anticoagulant | 3 minutes | With balanced heparin anticoagulant or lithium heparin if labeled for the measurement of electrolytes (for pH, PCO2, and PO2 only) *Not applicable for Ionized Calcium | 3 minutes *Not applicable for Ionized Calcium
 | With balanced heparin anticoagulant (or lithium heparin anticoagulant for pH, PCO2, and PO2 only) (syringe must be filled per manufacturer's recommendation) Maintain anaerobic conditions. Remix thoroughly before filling cartridge. | 10 minutes |  | With lithium heparin anticoagulant (tubes must be filled per manufacturer's recommendation) Maintain anaerobic conditions. Remix thoroughly before filling cartridge. | 10 minutes |  | 

## Slide 27: Prepare for sample collection - cont’d
BLOOD COLLECTION OPTIONS (part 2 of 2)
Analyte | SYRINGES | Test Timing |  | EVACUATED TUBES | Test Timing | CAPILLARY* | Test Timing
Sodium, Potassium, Hematocrit | Without anticoagulant | 3 minutes |  | Without anticoagulant | 3 minutes | With balanced heparin anticoagulant or lithium heparin if labeled for the measurement of electrolytes *Not applicable for Potassium | 3 minutes *Not applicable for Potassium
 | With balanced heparin anticoagulant or lithium heparin anticoagulant (syringe must be filled per manufacturer's recommendation) Remix thoroughly before filling cartridge | 30 minutes Note: Do not report pH, PCO2, PO2, and ionized calcium results for samples run beyond 10 minutes from collection |  | With lithium heparin anticoagulant (tubes must be filled per manufacturer's recommendation) Remix thoroughly before filling cartridge | 30 minutes Note: Do not report pH, PCO2, PO2, and ionized calcium results for samples run beyond 10 minutes from collection |  | 

## Slide 28: Running the test
REMOVE THE CARTRIDGE FROM
THE CARTRIDGE POUCH– EG6+
Place the cartridge on a clean, flat, level, and vibration-free surface
Handle the cartridge by the sides
PRECAUTIONS
DO NOT touch the sensors on the top of the cartridge.
DO NOT apply pressure to the central area of the label because the calibration fluid pack inside could burst prematurely.
DO NOT expose cartridges to temperatures above 30°C (86°F).
DO NOT place the cartridge and/or analyzer in an oxygen-enriched atmosphere.
DO NOT
TOUCH THE SENSORS
DO NOT TOUCH THE CENTER OF THE CARTRIDGE

## Slide 29: Running the test
REMOVE THE CARTRIDGE FROM
THE CARTRIDGE POUCH – EG7+
Place the cartridge on a clean, flat, level, and vibration-free surface
Handle the cartridge by the sides
PRECAUTIONS
DO NOT touch the sensors on the top of the cartridge.
DO NOT apply pressure to the central area of the label because the calibration fluid pack inside could burst prematurely.
DO NOT expose cartridges to temperatures above 30°C (86°F).
DO NOT place the cartridge and/or analyzer in an oxygen-enriched atmosphere.
DO NOT
TOUCH THE SENSORS
DO NOT TOUCH THE CENTER OF THE CARTRIDGE

## Slide 30: Running the test – cont’d
WHOLE BLOOD: COLLECT AND MIX THE SAMPLE
Collect a whole blood sample into a lithium heparin evacuated blood collection tube
Invert the lithium heparin collection tube at least 10 times
If sample was collected into a syringe, invert syringe for 5 seconds, then roll between the palms (hands parallel to the ground) for 5 seconds, flip and roll for an additional 5 seconds
Expelling 2 drops before filling a cartridge is desired
PRECAUTIONS
Caution should be used when handling materials that may contain transmissible infectious agents.
Fill evacuated tube per manufacturer’s recommendations.
10x
5sec
5sec
5sec

## Slide 31: Running the test – cont’d
WHOLE BLOOD: FILL THE i-STAT CARTRIDGE
Fill the cartridge immediately after mixing
Direct the hub of syringe or tip of the transfer device (capillary* tube, pipette, or dispensing tip) into the sample well of the cartridge
Slowly dispense sample into the sample well until the sample reaches the fill mark indicated on the cartridge. Cartridge is properly filled when the sample reaches the ‘fill to’ mark a small amount of sample is in the sample well.
CORRECTLY FILLED TO FILL LINE
IMPORTANT
*Capillary samples not applicable for Potassium and Ionized Calcium tests.

## Slide 32: Running the test – cont’d
CLOSE AND SEAL THE i-STAT CARTRIDGE
Fold the snap closure over the sample well to close and seal the cartridge
PRECAUTIONS
Incorrect handling and filling of the i-STAT cartridge will result in a quality check failure.
    CORRECTLY FILLED CARTRIDGE
    UNDER-FILLED CARTRIDGE
    OVER-FILLED CARTRIDGE

## Slide 33: Running the test – cont’d
INSERT THE CARTRIDGE
Insert the cartridge into the i-STAT analyzer’s cartridge port until it clicks into place.
  Sensors facing up, and on leading edge
PRECAUTIONS
DO NOT touch the sensors on the top of the cartridge.
DO NOT apply pressure to the central area of the label because the calibration fluid pack inside could burst prematurely.
DO NOT place the cartridge and/or analyzer in an oxygen-enriched atmosphere.

## Slide 34: Running the test – cont’d
PRIOR TO TESTING
The analyzer will perform quality checks
Failure of a quality check will result in a Quality Check Code

## Slide 35: Running the test – cont’d
TIME TO RESULTS
Once the cartridge is inserted, Contacting Cartridge will display followed by the countdown bar
The i-STAT EG6+ & i-STAT EG7+ cartridges each take approximately 130-200 seconds
This allows the user to estimate the time to results
Alerts such as “Cartridge Locked” are also displayed
PRECAUTIONS
DO NOT attempt to remove the cartridge while the message “Cartridge Locked” remains on the screen.
DO NOT REMOVE CARTRIDGE
Time To
Results
DO NOT REMOVE CARTRIDGE

## Slide 36: Running the test – cont’d
TIME TO RESULTS
Once the cartridge is inserted, Contacting Cartridge will display followed by the countdown bar
The i-STAT EG6+ & i-STAT EG7+ cartridges each take approximately 130-200 seconds
This allows the user to estimate the time to results
Alerts such as “Cartridge Locked” are also displayed
PRECAUTIONS
DO NOT attempt to remove the cartridge while the message “Cartridge Locked” remains on the screen.
Time To
Results
DO NOT REMOVE CARTRIDGE

## Slide 37
PERFORMING A PATIENT TEST
Avoiding Errors

## Slide 38: Troubleshooting basic errors
TROUBLESHOOTING & ERRORS
ERROR MESSAGES AND QUALITY CHECK CODES
There are potential pre-analytical conditions, especially during the training period, that may not be detected by the operator.
The analyzer will detect the condition, halt the test cycle, and display a cause message followed by the action message “USE ANOTHER CARTRIDGE.”
QUICK TIPS TO PREVENT ERRORS
The cartridge and analyzer must be at room temperature (18-30 °C or 64-86 °F) to perform the test.
The cartridge must be correctly filled, closed and sealed.
The analyzer must remain on a level surface. Level surface includes performing the test with the analyzer in the downloader/recharger.

## Slide 39: Factors affecting Sodium results
AVOID AFFECTING SODIUM RESULTS
  Draw adequate waste prior to drawing blood
  Avoid drawing blood from an arm with an IV line
IMPORTANT
Mix thoroughly and test promptly

## Slide 40: Factors affecting Potassium2 results
AVOID AFFECTING POTASSIUM RESULTS
  Prevent hemolysis resulting from traumatic draws / vigorous sample handling
  Do NOT ice samples
  Test promptly
  Avoid extra muscular activity drawing blood
IMPORTANT
Mix thoroughly and test promptly
Capillary sample types not applicable for Potassium tests.
TEST PROMPTLY
HEMOLYSIS

## Slide 41: Factors affecting Hematocrit results
AVOID AFFECTING HEMATOCRIT RESULTS
  Mix thoroughly
  Draw adequate waste prior to drawing blood
  Keep handheld level during testing
IMPORTANT
Mix thoroughly and test promptly
10x
5x

## Slide 42: Factors affecting pH results
AVOID AFFECTING PH RESULTS
  Avoid extra muscular activity during sample collection
  Do NOT expose to air
IMPORTANT
Mix thoroughly and test promptly

## Slide 43: Factors affecting PCO2 results
AVOID AFFECTING PCO2 RESULTS
  Test promptly
  Fill blood collection device completely
  Do NOT expose to air
IMPORTANT
Mix thoroughly and test promptly

## Slide 44: Factors affecting PO2 results
AVOID AFFECTING PO2 RESULTS
  Test promptly
  Do NOT ice samples
  Fill blood collection device completely
  Do NOT expose to air
IMPORTANT
Mix thoroughly and test promptly
TEST PROMPTLY

## Slide 45: Factors affecting Ionized Calcium2 results on EG7+
AVOID AFFECTING IONIZED CA RESULTS | EG7+
  Fill blood collection devices completely
  Limit extra muscular activity during sample collection
IMPORTANT
Mix thoroughly and test promptly
Capillary sample types not applicable for Ionized Calcium tests.

## Slide 46
Congratulations!
You have completed the lesson
i-STAT 1 EG6+ & EG7+
Cartridges USA Only