## Slide 1: Please delete this slide after reading through and before working on your slides
TIPS
Content provided in this lesson does not replace the i-STAT Alinity Quick Reference Guide, i-STAT Alinity System Operations Manual, or Instructions for Use (IFU) documents
Refer to the Instructions for Use (IFU) documentation for specific guidance on cartridge handling and sample collection.
This slide deck may be modified to represent products within the i-STAT System currently in use within your facility.
Once the deck has been modified, customers will assume responsibility for the content.*
Abbott will continue to provide and support the original resources that are made available on the www.globalpointofcare.abbott website.*
Not all products are available in all regions.
*Check the website regularly to ensure that you have the most up-to-date content.

## Slide 2: Cartridge Handling & Sample Collection
i-STAT Alinity System

## Slide 3: Learning Objectives
Describe the basic cartridge handling guidelines
Describe basic sample handling
UNDERSTAND THE FOLLOWING:

## Slide 4
Cartridge Handling
i-STAT ALINITY CARTRIDGE HANDLING & SAMPLE COLLECTION

## Slide 5: Anatomy of a Cartridge Box
Refrigerated storage temperature indicator: 2-8° C (35-46° F)
Indicates shelf life when stored at room temperature
Refrigerated storage expiration date
Cartridge LOT number
Location to record room temperature expiration date
Cartridge List Number
A
B
C
D
E
F
A
B
C
D
E
F

## Slide 6: Anatomy of a Pouch | Front
Cartridge name
Analyte(s) measured & calculated, if applicable
Location to record room temperature expiration date
2D barcode for manufacturing quality control; not scannable
Cartridge LOT number
Cartridge pouch barcode
Refrigerated storage expiration date
A
B
C
D
E
F
G

## Slide 7: Anatomy of a Pouch | Back
CARTRIDGES CAN BE STORED AT ROOM TEMPERATURE
Once cartridge is out of refrigerator for more than 5 minutes it should not be re-refrigerated, and the expiration date changes
Room temperature expiration date cannot exceed manufacturer’s printed expiration date
Indicates shelf life when stored at room temperature 14d = 14 days
Indicates shelf life when stored at room temperature 2m = 2 months
Room temperature storage range: 18°- 30° C (64°- 86°F)
A
B
C
C
A
B
C

## Slide 8: Anatomy of a Pouch | Back - cont’d
CARTRIDGES MUST BE BROUGHT TO ROOM TEMPERATURE BEFORE USE
Individual cartridge pouches must be at room temperature for 5 minutes before use
A box of cartridges must be at room temperature for 1 hour before use
5 mins
60 mins
IMPORTANT
Date cartridge pouches and cartridge boxes with the room temperature expiration date.
Do NOT manipulate the cartridge in any way to bring it to room temperature more quickly.

## Slide 9: Anatomy of a Portion Pack
FILL THE CARTRIDGE
Cartridge name
Analytes – measured & calculated, if applicable
2D barcode for manufacturing quality control; not scannable by the instrument Cartridge LOT number
Cartridge portion pack barcode for scanning
Refrigerated storage expiration date
Refrigerated-temperature-storage range
A
B
C
D
E
F
G
Portion Pack Front
Portion Pack Back
A
B
C
D
E
F
G

## Slide 10: Anatomy of a Cartridge
A
A
B
C
D
E
Contact Pads & Sensors (do not touch)
Calibrant pack (do not touch)
Closure
Fill to mark
Sample chamber
Sample well
F
B
C
D
E
F
A
B
C
F
D
E

## Slide 11: Cartridge Overview
i-STAT CARTRIDGES
Analysis, sample manipulation, and waste are confined within the cartridge
Dispose of used cartridges as biohazardous waste; follow facility policy for disposal

## Slide 12: i-STAT Cartridge Handling
IMPORTANT To prevent a Quality Check Code, avoid touching the sensors or pressing the center of the cartridge.
Carefully tear open the cartridge pouch or portion pack where indicated
Hold the cartridge by the sides or the bottom
DOs
Do not touch the sensors
Do not apply pressure to the center of the cartridge
DON'Ts

## Slide 13
Sample Collection
i-STAT CARTRIDGE HANDLING & SAMPLE COLLECTION
For in vitro diagnostic use only. For intended use and product details, visit www.globalpointofcare.abbott. Not all products are available in all regions. Check with your local representative for availability.
Photos & images are for illustrative purposes only. ©2025 Abbott. All rights reserved. i-STAT & Alinity are a trademarks of Abbott.| iLS Presentation i-STAT Alinity – Cartridge Handling and Sample Collection | APOC-25002608.2

## Slide 14: Sample Collection Techniques
i-STAT ALINITY SYSTEM CAN BE USED WITH THE FOLLOWING SAMPLE TYPES:
IMPORTANT Follow your hospitals procedures for proper sample collection. Correlate the results with the patient’s status.
Arterial | Venous | Capillary
Sample type varies based upon cartridge type. See Instructions for Use (IFU) documents on the APOC website www.globalpointofcare.abbott to determine the appropriate sample type for the cartridge being used

## Slide 15: Sample Collection | Evacuated Tube
ENSURE EVACUATED TUBES ARE FILLED ACCORDING TO THE TUBE MANUFACTURER’S RECOMMENDATION.

## Slide 16
MIXING MATTERS!
Mix sample by gently inverting tube at least 10 times!
Sample Mixing | Evacuated Tube
10x

## Slide 17
Mixing matters!
Always mix properly!
Improper mixing may affect results
Sample Handling | Evacuated Tube

## Slide 18
DO NOT ice sample
DO NOT allow air bubbles to be trapped in the syringe
DO NOT fill the cartridge by dripping blood into the sample well
Sample Handling | Syringes

## Slide 19
Sample Mixing | Syringes
MIXING MATTERS!
Mix sample by vigorously rolling the syringe between palms for 5 seconds; invert and repeat.
5 SEC
5 SEC

## Slide 20
Sample Handling | Syringes
MIXING MATTERS!
Always mix properly!
Improper mixing may affect results

## Slide 21
CAUTION
Ensure air bubbles are not trapped in the capillary tube
Sample Collection | Capillary

## Slide 22
Use a skin puncture device that provides free-flowing blood
Samples collected into a capillary tube must be tested immediately
Sample Handling | Capillary

## Slide 23
Filling a Cartridge
SAMPLE WELL
Fill to mark
SAMPLE WELL
FILL TO MARK
Fill to mark
SAMPLE WELL

## Slide 24
Filling and Closing the Cartridge - cont’d
Place the tip of the transfer device into the sample well as shown
Fill cartridge until sample reaches fill mark as shown
Fold over snap closure or slide until a soft click is heard and closure is secured
1
2
3

## Slide 25
Avoid Cartridge-filling Quality Checks
INCORRECT FILL
Overfilled, Underfilled, Improperly Closed and Air Bubble
CORRECT FILL
Filled to fill mark and securely closed.

## Slide 26
Level-Surface Testing
ADDITIONAL REQUIREMENTS:
Indoors, on a dry, clean, horizontal, stable surface.
Avoid nearby vibration equipment such as centrifuges.
For ACT, HcT, GFAP, UCH-L1, and hs-Tnl testing, the analyzer must remain on a level surface with the display facing up during testing. A level surface includes running tests on the instrument with it resting in the Base Station.

## Slide 27: Factors that may affect results
  Avoid factors that may affect the sample prior to analysis
  Always correlate results with the status of the patient
  For additional information please visit www.globalpointofcare.abbott
For a comprehensive list of factors that may affect results, go to Abbott’s Global Point of Care website for product Instructions for Use (IFU) and other technical documentation

## Slide 28: Summary
Describe the basic cartridge handling guidelines
Describe basic sample handling
NOW THAT YOU HAVE COMPLETED THIS LESSON, YOU SHOULD BE ABLE TO:

## Slide 29: Summary
Congratulations!
You have completed the lesson i-STAT Alinity - Cartridge Handling & Sample Collection